Summary of Calibration Factors

C-130 - Basic ICE-T Configuration

Heymsfield et al

Pre-Calibration – 06/20/2011

Param

Variable                    Name      C1        C2       C3

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   1.042	  17.377  -.0060

Dynamic Pressure-Radome (hPa)	QCR 	  -1.163	  17.415	-.0023

Dynamic Pressure-Fuselage R. (hPa)	QCFR	  -0.206	  13.916  -.0026

Total Temperature-Radome Left (oC)	TTRL	 -67.954	  27.090	-.4603

Total Temperature 1:Heated Fuse(oC)  TTHR1   -93.032    28.854  -.0988

Total Temperature 2:Heated Fuse(oC)  TTHR2   -93.418    28.699  -.0911

Remote Surface Temp. Bot. (oC) 	RSTB    -46.563     8.131   .1631

Remote Surface Temp.-Bot. (oC) 	RSTB1   -34.585     2.796   .6279

Remote Surface Temp. Top. (oC)	RSTT    -50.769     9.528   .0702

Vertical Press. Diff. Radome(mb)	ADIFR	   0.028     7.020	 .0091

Horizontal Press. Diff. Radome(mb) 	BDIFR	   0.001     5.179   .0006

CVI Inlet Pressure (hPa)             CVPCN    -5.817   342.050  -.0354

CN Flow Rate (slpm)	FCN	   0.000	   0.323   .0000

Dew Point Sensor Bottom (C)	DPB	  -1.5       1.0     .0

Dew Point Sensor Top (C)	DPT	  -1.5       1.0     .0

Post Calibration – 08/08/2011

Param

Variable                    Name      C1        C2       C3__

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.962	  17.265   .0003

Dynamic Pressure-Radome (hPa)	QCR 	  -1.163	  17.376	-.0057

Dynamic Pressure-Fuselage R. (hPa)	QCFR	  -0.610    13.848  -.0048

Total Temperature-Radome Left (oC)	TTRL	 -68.156	  27.110	-.4666

Total Temperature 1:Heated Fuse(oC)  TTHR1   -93.141    28.932  -.1116

Total Temperature 2:Heated Fuse(oC)  TTHR2   -93.225    28.570  -.0724

Remote Surface Temp. Bot. (oC) 	RSTB*#  -59.260    11.067   .0000

Remote Surface Temp.-Bot. (oC) 	RSTB1*# -62.207    11.422   .0000

Remote Surface Temp. Top. (oC)	RSTT    -54.692    10.638  -.0101

Vertical Press. Diff. Radome(mb)	ADIFR	  -0.043     6.978	-.0014

Horizontal Press. Diff. Radome(mb) 	BDIFR	  -0.040     5.149   .0004

Cabin Pressure (hPa)                 PCAB     -1.778   109.350  -.0380

CN Inlet Pressure (hPa)              PCN#     -8.866   108.780  -.0333

CN Flow Rate (slpm)	FCN#	   0.000	   0.452   .0000

Dew Point Sensor Bottom (C)	DPB	  -0.400     1.035   .0018

Dew Point Sensor Top (C)	DPT	  -0.597     1.059   .0038

# Post Cal used in production processing

* Water Calibration

Calibration Summary

page 2

Calibration Maneuvers –

RAF uses sets of in-flight maneuvers specifically designed to help evaluate the performance of the 3-D wind measurement system on the aircraft.  Sets of maneuvers were flown on one test flight and two research flights: TF01 and RF11 respectively. In each case the variations observed fell within acceptable limits.

PMS - 1-D Probe Bin Sizes

The SPP100 (std), SPP200, CDP and SPP300 probes were used on the project.  The minimum particle size (microns) for each sampling bin are listed below:

SPP100

Std  		SPP200           CDP  		SPP300

bin#1 		 1.1   um		 0.097 um		  2.25 um		0.266

bin#2		 	 2.51		 0.105		  3.25		0.284

bin#3			 3.92 		 0.113		  4.25		0.32

bin#4			 5.33		 0.121		  5.25		0.371

bin#5			 6.74		 0.129		  6.25		0.423

bin#6			 8.15		 0.145		  7.25		0.475

bin#7			 9.56		 0.162		  8.25		0.528

bin#8			10.97		 0.182		  9.25		0.581

bin#9			12.38		 0.202		 10.25		0.635

bin#10		13.79		 0.222		 11.25		0.7

bin#11		15.20		 0.242		 12.25		1.06

bin#12		16.61		 0.262		 13.25 		1.795

bin#13		18.02		 0.282		 14.25		2.43

bin#14		19.43		 0.302		 16.125		3.449

bin#15		20.84		 0.401		 18.0		4.112

bin#16		22.25		 0.570	  	 19.875		4.792

bin#17		23.66		 0.656		 21.75		5.578

bin#18		25.07		 0.740	  	 23.625 		6.668

bin#19		26.48		 0.833		 25.5		7.728

bin#20		27.89		 0.917		 27.375		8.717

bin#21		29.30		 1.008		 29.25		9.665

bin#22		30.71		 1.148		 31.125	    10.651

bin#23		32.12		 1.319		 33.0	    11.636

bin#24		33.53		 1.479		 34.875	    12.623

bin#25		34.94		 1.636		 36.75	    13.614

bin#26		36.34		 1.796		 38.625	    14.647

bin#27		37.76		 1.955		 40.5	    15.595

bin#28		39.17		 2.184		 42.375	    16.683

bin#29		40.58		 2.413		 44.25	    17.776

bin#30		41.99		 2.661		 46.125	    18.876

bin#31		43.40		 2.991		 48.0	    20.0