## Sheet1
Date | Instrument | Issue | Solution | Flight # | Tech | Description - From Tech Notes File
2018-06-18 | Router | Installed new router and collapsed networks |  | Pre | JC | Installed new router in ADS rack and went to 1 network for the aircraft. The Satcom is not hooked directly to the new router and the router in the video rack is unused. There will not no icon for satcom on the desk top. all controlled by software now.
2018-06-18 | ADS | Not running on UPS for this project |  | Pre | JC
2018-06-18 | TimeServer | Not running on UPS for this project |  | Pre | JC
2018-06-18 | A/D AutoCal | Ran pre-project See description In Tech notes |  | Pre | JC
2018-07-13 | A/D AutoCal | Ran Preflight A/D Cal, see notes for details |  | TF01 | BI
2018-07-13 | TCAB | L-1 and R-8 not working correctly |  | TF01 | BI
2018-07-13 | SDI | SDI flow is not accurate |  | TF01 | BI
2018-07-13 | UFT4 | Under floor temp not accurate |  | TF01 | BI
2018-07-13 | UFT4 | Bay 3 is hottest at 113F |  | TF01 | BI
2018-07-13 | SDI | Not turned on until after takeoff |  | TF01 | BI
2018-07-13 | 2DC-10 | No voltage from probe |  | TF01 | BI
2018-07-13 | 2DC-10 | Unable to diaply images |  | TF01 | BI
2018-07-13 | SatComm | Lost connection at 1122L |  | TF01 | BI
2018-07-13 | UFT4 | Bay 3 now at 114F 1140L |  | TF01 | BI
2018-07-13 | TCAB | was at 90F now about 85F |  | TF01 | BI
2018-07-13 | SDI | Not enough flow for users, needs work |  | TF01 | BI
2018-07-13 | TCAB | L-6 now at 82F 1210L |  | TF01 | BI
2018-07-13 | Headset | Seat at L-7 Headset not working |  | TF01 | BI
2018-07-13 | SatComm | Lost again at 1212L back at 1228L |  | TF01 | BI
2018-07-13 | UFT4 | Bay 3 dropped to 97F at 20K feet |  | TF01 | BI
2018-07-13 | SatComm | Lost again at 1228L |  | TF01 | BI
2018-07-13 | HGM232 | Aeros is NAN local display okay |  | TF01 | BI
2018-07-16 | SPP200 | Would not connect with either RS232 or 422
2018-07-16 | A/D AutoCal | Ran today. |  |  |  | same as before except DSM 323 card 0 ch0 showing 20mv offset.
2018-07-16 | TCAB | L-1 looking at why reading is weird in Aeros
2018-07-17 | ADS server | Software change applied then backed out preflight |  | TF02 | DA
2018-07-17 | Power Shift | Ground to APU had problems |  | TF02 | DA | Power transfer (GND to APU) was attempted at 39.5 A on the Left Cabin Bus. This was unsuccessful as some equipment had still lost power during the transfer.
2018-07-17 | SPP200 | Not on for this flight |  | TF02 | DA
2018-07-17 | Cloud Coll | Not on for this flight |  | TF02 | DA
2018-07-17 | 2DC-10 | 2D images do not display correctly |  | TF02 | DA
2018-07-17 | SatComm | Various dropouts once for 20 minutes |  | TF02 | DA | Several SATCOM dropouts. Once out for ~ 20 minutes, cycled the switch in cockpit (1932), and only came back for 60 seconds or so.
2018-07-17 | VNC | VNC icon for VCSEL did not work correctly |  | TF02 | DA
2019-07-18 | SPP200 | Tested serial interface, now Ser 9 | Ser 5, 6 and maybe 11 bad |  |  | Did testing on PCASP. with DMT software proved out that can wiring was good. Went to Ser 5 and it did not work. Tried loop back and it did not work. Went to Ser 9 and it worked. Left probe on Ser 9 for the project.
2018-07-19 | SatComm | Tested SatComm | Replaced RF Switch |  |  | Satcom testing. Towed plane in a circle and found that when changed to Starboard antenna (I thin) it would try and dial and AFIS sat that would could not do internet with. Also suspected that the RF switch might be bad. RF switch replaced.
2018-07-20 | SatComm | No dropouts during flight |  | FF01 | DA | SATCOM stayed up for the entire 2.5 hr flight, Craig Poyner confirmed via logs. A 360 circle maneuver was done @ ~ 17:50 (11:50) MST the connection stayed up. Prior to departure on 7/19 the high power relay was swapped for a new unit, P/N 301706-500 . At this point more testing needs to be completed to verify if the swap truly fixed the antenna hand-off issue. KAZ
2018-07-21 | 2DC-10 | Probe worked on
2018-07-23 | General | RF 01 Flight scrubbed
2018-07-23 | RTH | Both elements went to NAN | Sun overheating them
2018-07-23 | DSM | RW DSM did not boot initially
2018-07-24 | A/D AutoCal | DSM 323 A2D 1606 20 mv difference |  | RF01 | DA
2018-07-24 | TCAB | L-1 intermitiant at turn on |  | RF01 | DA | L1 TCAB intermittently cycling ON and OFF at initial power ON. Opened up box, checked battery, wiggled wires and connections…could not recreate failure.
2018-07-24 | ADS server | Another update that had to be removed. |  | RF01 | DA
2018-07-24 | DSM | RW DSM did not boot initially |  | RF01 | DA
2018-07-24 | SMPS | Tripped braker twice and turned off |  | RF01 | DA
2018-07-24 | SatComm | 2 dropouts of about 2 minutes each according to xchat |  | RF01 | DA
2018-07-25 | PLWC | Looked at probe |  |  |  | Measured PLWC element, it measured correctly. Cleaned the element connectors pins and sockets as best as possible. Measured contacts again. Re-installed
2018-07-25 | RSTB | Cleaned windows of both instruments
2018-07-25 | VCSEL | Cleaned optics
2018-07-25 | CDP | Did Calibration
2018-07-25 | ICS | Headset cord R-1 Broken
2018-07-26 | A/D AutoCal | Completed Preflight |  | RF02 | BI
2018-07-26 | Cameras | Stopped working SE to look at |  | RF02 | BI
2018-07-26 | RSTB | Julie qustioning data, did not see onboard |  | RF02 | BI
2018-07-26 | Probes | CDP and SPP200 tracking nicely |  | RF02 | BI
2018-07-27 | CDP | Cleaned wiindows and calibrated
2018-07-27 | DSM | RW Pod DSM continues to not want to boot. |  |  |  | RW pod DSM gave trouble this morning, reset all connectors, including power, used power switch on DSM...and it started working. Rebooted it multiple times, worked every time here tonight.
2018-07-27 | VCSEL | Cleaned mirrors |  |  |  | Kurt worked on VCSEL using camera paper to clean...best he got was about 1.7-1.8 volts of LSRINTAMB
2018-07-27 | PAN | Inlet removed
2018-07-27 | 2DC-10 | Probe Removed for work
2018-07-28 | General | RF03 canclled due to timeserver problems
2018-07-30 | Timeserver | Timeserver in simulation mode for this flight |  | RF03 | DA | TimeServer in Simulation Mode for this flight. Connected to zero (0) satellites but has good time. Will test in normal mode post flight. *Normal mode results failed – 0 satellites connected.
2018-07-30 | QCR | QCR Iced up |  | RF03 | DA
2018-07-30 | RSTB | Displays are choppy with Intermittant data | Sensor maxed out at 55C | RF03 | DA | Sensors max out at 55C ground temp
2018-07-30 | Probes General | All probes seem to be working normally |  | RF03 | DA
2018-07-31 | Timeserver | Timeserver in simulation mode for this flight |  | RF04 | DA | TimeServer in Simulation Mode for this flight. Put in Normal Mod pre-flight, saw 9 out of 13 satellites.
2018-07-31 | UHSAS | Flow set at 10 per Reeves |  | RF04 | DA
2018-07-31 | RSTB | One drop after takeoff |  | RF04 | DA
2018-07-31 | QCR | Worked okay even after icing last flight |  | RF04 | DA
2018-07-31 | Probes General | Lots of probe activity this flight |  | RF04 | DA
2018-08-01 | TimeServer | Back to Normal mode
2018-08-01 | VCSEL | Cleaned VCSEL
2018-08-01 | CDP | Calibrations and Cleaned windows
2018-08-01 | Dew Point | Cleaned
2018-08-01 | UVHygrometer | Cleaned
2018-08-01 | Radome | Drained lines
2018-08-01 | 2DC Probes | Cleaned Mirrors
2018-08-01 | RTF1 | Checked probe for damage | No damage found |  |  | Fast temp looked for damaged element found no issues. Reinstalled
2018-08-02 | TimeServer | In Normal mode |  | RF05 | DA
2018-08-02 | DSM | DSM 319 had an issue | Power cycled and cam back up | RF05 | DA | 1826 – Lost DSM 319; cycled power, came right back, no further issues
2018-08-02 | RSTB | Started dropping data @ 1927 | Each drop occurred at about 55C | RF05 | DA
2018-08-02 | UHSAS | Probe Stopped @0018 |  | RF05 | DA
2018-08-02 | VCSEL | Laser Power dropped thoughout the flight |  | RF05 | DA
2018-08-03 | VCSEL | Cleaned before RF06 and repalced thumb drive |  | RF06 | BI | Drive crashed and not being used any more
2018-08-03 | Temp Probes | Look good after Takeoff |  | RF06 | BI
2018-08-03 | Dew Point | DPXC not working |  | RF06 | BI
2018-08-03 | DSM | DSM 319 restatred by SE on ground @ 1506L |  | RF06 | BI
2018-08-03 | UHSAS | 2 laser resets during the flight 1600L |  | RF06 | BI
2018-08-03 | SatComm | Many Satcomm failures |  | RF06 | BI | Many Sat Comm failures...on/off, some every 3-6 minutes Total so far 8 times for 3-9 minutes
2018-08-03 | VCSEL | Soft boot at 1830L Fixed problem |  | RF06 | BI
2018-08-04 | VCSEL | Cleaned, Data downloaded
2018-08-04 | CDP | Cleaned and Calibrated
2018-08-06 | Timeserver | Timeserver in simulation mode for this flight |  | RF07 | DA | See tech notes for detailed description
2018-08-06 | DSM | DSM 319 power cycled, pressures NAN |  | RF07 | DA
2018-08-06 | RTF1 | Probe changed prior to flight | Old S/N 3243, NEW S/N 2932 | RF07 | DA
2018-08-06 | RSTB | Data goes to NAN when inputs are above 55C |  | RF07 | DA
2018-08-06 | Cameras | Stopped working at 2300 |  | RF07 | DA
2018-08-06 | UHSAS | Laser cycled at 2355 |  | RF07 | DA
2018-08-06 | Probes General | All working okay |  | RF07 | DA | UHSAS, CDP and PCASP all functioning normally,2D images occasionally when in heavier smoke, larger aerosols
2018-08-07 | Timeserver | Worked on it again |  |  |  | See tech notes for detailed description
2018-08-07 | VCSEL | Cleaned
2018-08-07 | CDP | Cleaned and Calibrated
2018-08-07 | ADS Server | Cleaned Firewire card contacts
2018-08-08 | TimeServer | In Normal mode |  | RF08 | DA
2018-08-08 | UFT4 | Low level flying Bay 3 was at 126F |  | RF08 | DA
2018-08-08 | TCAB | R-6 reached 90F |  | RF08 | DA
2018-08-08 | DSM | DSM 319 power cycled, pressures NAN @1858 |  | RF08 | DA
2018-08-08 | DSM | DSM 319 power cycled, pressures NAN @0002 |  | RF08 | DA | 0002 – Lost QCF, ADIFR, BDIFR and DPB. *(DPB is on DSM 321) Cycled power on only DSM 319 and all variables came back
2018-08-08 | SatComm | Down all Pre-flight, came back for flight |  | RF08 | DA | SATCOM was down all of pre-flight. It was decided that SATCOM was not essential to the research mission. Had dial-tone on the handset in Video Rack but could not ping internet or connect via Xchat. Cycled switch in cockpit once (shortly after takeoff) and SATCOM came up within 10 minutes.
2018-08-09 | VCSEL | Cleaned Pre-flight |  | RF09 | DA
2018-08-09 | CDP | Cleaned Preflight |  | RF09 | DA
2018-08-09 | RSTB | Cleaned windows Preflight, RSTB1 pitted |  | RF09 | DA
2018-08-09 | Timeserver | Timeserver is in Normal for this flight |  | RF09 | DA
2018-08-09 | DSM | DSM 325 needed to have a power cycle @ startup |  | RF09 | DA
2018-08-09 | Probes General | All probes looked good on the flight |  | RF09 | DA
2018-08-10 | CDP | Cleaned and Calibrated, Times in TechNotes
2018-08-10 | UVHygrometer | Cleaned
2018-08-10 | Dew Point | Both Cleaned
2018-08-10 | RTF1 | Cleaned and heat cover made.
2018-08-10 | DSM | DSM 319 removed and cleaned |  |  |  | Removed DSM 319, took apart and cleaned with contact cleaner. Reinstalled. Josh saw that the timing card was resetting and the analog card was resetting. Hoping cleaning will fix it. Noticed a pretty heavy coat of dust on some cards during cleaning
2018-08-11 |  | Media and open house Day
2018-08-13 | Timeserver | Timeserver in simulation mode for this flight |  | RF10 | KH
2018-08-13 | VCSEL | Cleaned Preflight |  | RF10 | KH
2018-08-13 | DSM | DSM 319 needed reboot after power dump |  | RF10 | KH
2018-08-13 | SatComm | needed rebot after power dump |  | RF10 | KH
2018-08-13 | QCR | Iced up @ 2338, radome heat on. |  | RF10 | KH
2018-08-13 | RTF1 | Icing may have slightly effected the probe |  | RF10 | KH
2018-08-14 | CDP | Cleaned and Calibrated CF 07 Times in Tech Notes
2018-08-14 | Timeserver | Timeserver troubleshooting see Tech Notes
2018-08-14 | DSM | DSM 319 Replaced Timing card | Card 40869 out, 30274 IN |  |  | See Tech notes for Troubleshooting information
2018-08-14 | 2DC-25 | End Diode voltages trouble shooting | Voltages saturate in flight |  |  | See Tech notes for Troubleshooting information
2018-08-14 | HARP | Replaced top radiometer gasket
2018-08-14 | Radome | Opened for water check, no water in the lines.
2018-08-15 | 2DC-25 | Replaced prior to flight |  | RF11 | JC | Replaced 2DC-25 S/N 3 with S/N 1. put on a spinning disk and the particles looked slightly squished. Looked at end dioses and they were 1.0 V. Decided to verify mirrors were clean. after cleaning 1.1, 2.98, 0.9. did not recheck with disk
2018-08-15 | Timeserver | Normal mode for flight |  | RF11 | JC
2018-08-15 | VCSEL | Cleaned before flight |  | RF11 | JC
2018-08-15 | A/D AutoCal | Ran before flight |  | RF11 | JC | Ran A/D Auto Cal. DSM 324 card 0 ch 3 RSTT is 8mv offset, DSM 323 card 0 ch0 UFT4 is 21 mv offset rest of the cards looked good
2018-08-15 | HGM232 | NAN on data system, head displays correctly. |  | RF11 | JC
2018-08-15 | UFT4 | Bay 4 107F 30 min after takeoff |  | RF11 | JC
2018-08-15 | DSM | DSM 323 getting No Code indiacation at 2045 | Cleared at 2055 | RF11 | JC
2018-08-15 | TRSTB | Breaker found off, turned on. |  | RF11 | JC
2018-08-15 | 2DC-25 | Saturating on diode 32 again |  | RF11 | JC
2018-08-15 | 2DC-25 | Looks like TAS reported by 2DCD is 2 MS faster than plane |  | RF11 | JC | 2131 Looks like the TAS reported by the 2DC-25 is about 2 m/s faster than what the plane is reporting. I noticed that the spinning disk reported 35 m/s and I thought it was supposed to be 33 m/s
2018-08-15 | UFT4 | 2251 Underfloor temps bay 1 to bay 6: 93 89 100 97 95 96 |  | RF11 | JC
2018-08-15 | DSM | DSM 319 serials messed up. DSM reset @ 0008 |  | RF11 | JC
2018-08-15 | Power | 400 Hz Meter at Power Box not reading correctly | Only on Engine power | RF11 | JC
2018-08-16 | CDP | Cleaned, No calibration |  | RF12 | TA
2018-08-16 | VCSEL | Cleaned and Temp Tip Broke off |  | RF12 | TA
2018-08-16 | Cameras | Required restart at beginning of flight |  | RF12 | TA
2018-08-16 | Probes General | UHSAS, CDP and PCASP worked in clouds |  | RF12 | TA
2018-08-17 | UVHygrometer | Cleaned
2018-08-17 | CDP | Cleaned and Calibrated
2018-08-17 | VCSEL | Cleaned
2018-08-17 | 2DC-10 | Adjusted it physically to get End Diodes lower | 32 End diode was 1.45V |  |  | Looked at the 2DC per Josh's e-mail. I noticed that as I loosened the back hemisphere that the end diode voltages decreased from about 2.8 down to 1.7V. We wiggled connections but could not find anything that looked like a bad or weak connections. There were no changes at all any ware on the connection string. When we were completed I loosened the mirror screws and wiggled them until we had a 32 end diode voltage of 1.45V. Everything is now tight and back together.
2018-08-19 | DSM | Booted all DSM's and confirmed they were working
2018-08-19 | A/D AutoCal | Did Calibration |  |  |  | Did an A2D calibration. DSM 323 A0, Channel 0, Serial #1606 is more than 10mV off from original calibration. It is the UFT4 variable. We think it's not necessary to change since this variable is not really reading correctly
2018-08-20 | Cameras | Took 4 times to get the cameras started |  | RF13 | KH
2018-08-20 | DSM | DSM 323 had No Code at start, cleared later |  | RF13 | KH
2018-08-20 | 2DC-10 | End diodes look to be about double what is expected. |  | RF13 | KH
2018-08-20 | 2DC-25 | Voltages look okay with out railing. |  | RF13 | KH
2018-08-20 | DSM | DSM 324 may have had channel shift rebooted |  | RF13 | KH | DSM 324 may have had an analog channel shift. Rebooted at 19:21 with an analog fix but the serial card was not happy. Rebooted again at 19:24 with everything looking good.
2018-08-20 | Power | 400Hz Freq off when on engine power |  | RF13 | KH
2018-08-20 | Probes General | Saw returns from all probes during flight |  | RF13 | KH
2018-08-20 | Dew Point | DPB less responsive then DPT (need cleaning?) |  | RF13 | KH
2018-08-20 | RTF1 | Might be getting gunked up due to smoke (clean)? |  | RF13 | KH
2018-08-23 | General | Plane was gone for 2 days due to maintenance issue |  | RF14 | JC
2018-08-23 | CDP | Cleaned windows |  | RF14 | JC
2018-08-23 | VCSEL | Cleaned |  | RF14 | JC
2018-08-23 | A/D AutoCal | Ran Calibration, RSTT 9 mV difference but seems stable. |  | RF14 | JC
2018-08-23 | 2DC-25 | End Diodes were 0.59 0.78 0.98 at startup |  | RF14 | JC
2018-08-23 | General | Aborted Takeoff due to Oil Pressure |  | RF14 | JC
2018-08-23 | Power | Went to APU 400 Hz Freq okay Only bad on Engines |  | RF14 | JC
2018-08-23 | General | Lost power at end of runway, restarting |  | RF14 | JC
2018-08-23 | Power | Data Bus Voltage meter reading low. | Verified bus so it is just the meter | RF14 | JC | 2309 The meter for Data Power voltage is reading strange. Was 95V. I tapped it and now it reads 58V. I have measured the voltage at a plug and it reads 115V. It looks like a meter issue and we will chase it down next maintenance day if we get a chance. If not we will just monitor it
2018-08-23 | 2DC-25 | End diodes looked good |  | RF14 | JC
2018-08-23 | UHSAS | at 0131 Turned off laser per Reeves |  | RF14 | JC
2018-08-23 | UFT4 | On Ferry bay 4 96, bay 2 76 others in mid 80's |  | RF14 | JC
2018-08-23 | Cloud Coll | When removed there was water inside the can |  | RF14 | JC
2018-08-24 | CDP | Calibrations and Cleaned windows
2018-08-24 | VCSEL | Cleaned
2018-08-24 | RTF1 | Flushed with water
2018-08-24 | Dew Point | Cleaned both DPT abd DPB
2018-08-24 | QCFR | Removed sall bug guts from probe |  |  |  | Removed some material that looked like small bug guts from QCFR after Britt and I looked at the probes. Will need to see if it improves the measurement
2018-08-25 | Cameras | Failed to start, coule not get started in flight |  | RF15 | TA
2018-08-25 | VCSEL | 2045 Restarted per Reeves |  | RF15 | TA
2018-08-25 | VCSEL | 2058 Restarted again per Reeves but would not work |  | RF15 | TA | 20 :58 Resstart VCSEL again by Mike Reeves but at this time VCSEL really broke No matter how hard I tried restart VCSEL program still not work, Therefore VCSEL down for the rest of this flight (Mike Reeves Confirm it's dead after we landed )
2018-08-25 | ADS | Rebooted server due to it being hung up. |  | RF15 | TA
2018-08-27 | VCSEL | Looked at it but could not fix. | Broke for rest of the project
2018-08-27 | CDP | Calibrations and Cleaned windows | More dirty then in the past.
2018-08-28 | Timeserver | Simulation mode this flight. |  | RF16 | KH | Timeserver came up in a Stratum 2 position. Tried all the usual fixes with no success. Forced it into Simulation Mode. Suspicion is interference from somewhere at the airport
2018-08-28 | ADS | DSM server crashed 3 times, Zombie process?
2018-08-28 | Novatel | Not seeing data on Aeros | Was linked to Timeserver issues
2018-08-28 | SatComm | Drop out at 1946
2018-08-29 | Timeserver | Came up in Stratum 2, left in Simulation Mode |  |  |  | We powered up the timeserver and it again came up in stratum 2. We took it to Simulation Mode and brought up the data system. It all came up fine today. We are going to leave it in Simulation mode for the flight back to Jeffco so we do not have to mess with it on a short preflight,
2018-08-31 | General | No issues on the flight |  | FF02 | JM
2018-09-06 | General | Education flight, not full instrumentation |  | RF17 | BI
2018-09-06 | 2DC | No displays in back except for ADS system |  | RF17 | BI
2018-09-10 | DSM | DSM 319 Serials messed up. | Needed to reboot all DSM's | RF18 | JM | Initial check out, no ADIFR, BDIFR or QCF. Cycled power DSM319, no change Cycled power all DSMs @1449hrs, problem fixed
2018-09-10 | A/D AutoCal | Did Calibration |  | RF18 | JM | A2D check: DSM318, A0 TRSTB (SN001, Ch0) 26mV off, | DSM321, A0 XCELLTEMP (SN1601, Ch2) 40mV off, DSM321, A0 XSIGV (SN1601, Ch3) 38mV off, DSM323, A0 UFT4 (SN1606, Ch0) 50mV off, will inform project mgt per instructions, No channel shifting noted
2018-09-10 | DSM | DSM 323 had no PPS to start. Cleared at 2149 |  | RF18 | JM
2018-09-13 | Timeserver | Stratum 2 initial, fixed before flight |  | RF19 | JM
2018-09-13 | ADS | Acserver did not start SE fixed |  | RF19 | JM
2018-09-13 | ADS | Network problem before flight | Switch needed new power cord | RF19 | JM
2018-09-13 | DSM | RTH1&2 both NAN | Rebooted DSM 318 | RF19 | JM
2018-09-13 | DSM | DSM 325 stopped | Rebooted DSM 325 | RF19 | JM
2018-09-13 | DSM | DSM 323 No Code at 1916, clears at 2129 |  | RF19 | JM
2018-09-13 | Dew Point | Cycled power on DPT at 2125 |  | RF19 | JM
2018-09-17 | CDP | Calibrations and Cleaned windows
2018-09-20 | A/D AutoCal | Final Calibration |  |  |  | Final A/D AuotCal Error message at end of running AutoCal A2D 002 (DSM 324) Ch 3 level -10V is -8.73406V Out of spec cards DSM 323 A2D 1606 Ch 0 has 20mV offset DSM 324 A2D 002 Ch 1 has 4 mV offset Ch 3 had 9 mV offset DSM 318 A2D 006 Ch 0 has 5 mV Offset