Summary of Calibration Factors

GV - Basic IceBridge Configuration

Studinger et al

Pre-Calibration - 9/29/11

Param

Variable                    Name      C1        C2       C3

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   1.126	  0.9966   .0

Dynamic Pressure-Radome (hPa)	QCR 	   0.16	  1.0025	 .0

Total Temp: Rosemont Heated #1 (C)	TTFH1	-101.3      30.254  -.4678

Total Temp: Rosemont Heated #2 (C)	TTFH2   -99.251    29.869  -.4849

Total Temp: HARCO Heated #1 (C)	TTHR1 	 -91.125    26.524  -.0881

Total Temp: HARCO Heated #2 (C)      TTHR2   -91.787    26.820  -.0967

Vertical Press. Diff. Radome(mb)	ADIFR	  -0.128     1.0022  .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.06      1.0014  .0

Static Pressure-Fuselage (hPa)	PSF	  -0.809     1.0014  .0

Dew Point Sensor Left Serial (C)	DPLS	  -1.5       1.0     .0

Dew Point Sensor Right Serial (C)	DPRS	  -1.5       1.0     .0

Cabin Pressure (hPa)	PCAB     -3.528   109.09  -.0159

Post Calibration - 11/08/11

Param

Variable                    Name      C1        C2       C3__

Dynamic Pressure-Fuselage L. (hPa)	QCF  	  -0.024	   1.001   .0

Dynamic Pressure-Radome (hPa)	QCR 	  -0.153	   1.001	 .0

Total Temp: Rosemont Heated #1 (oC)	TTFH1	-100.96     30.212  -.4635

Total Temp: Rosemont Heated #2 (C)	TTFH2   -98.833    29.756  -.4597

Total Temp: HARCO Heated #1 (oC)	TTHR1 	 -92.074    27.105  -.1679

Total Temp: HARCO Heated #2 (C)      TTHR2

Vertical Press. Diff. Radome(mb)	ADIFR	   0.048     0.9974  .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.114     0.9978  .0

Static Pressure-Fuselage (hPa)	PSF	  -0.024     1.0016  .0

Dew Point Sensor Left Serial (C)	DPLS

Dew Point Sensor Right Serial (C)	DPRS

Cabin Pressure (hPa)	PCAB     -3.304   109.52   -.0201

# Post cal used in final processing.

Calibration Maneuvers -

RAF uses sets of in-flight maneuvers specifically designed to help evaluate the performance of the 3-D wind measurement system on the aircraft.  Sets of maneuvers were flown on one test flight and one research flight: XXX respectively. In each case the variations observed fell within acceptable limits.