Summary of Calibration Factors

GV - Basic ADELE/SPRITE Configuration

Smith / Nielsen

Pre-Calibration – 7/31/09

Param

Variable                    Name      C1        C2       C3

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.345	  1.0011   .0

Dynamic Pressure-Radome (hPa)	QCR 	  -0.248	  1.007	 .0

Total Temperature-Radome Left (oC)	TTFR	 -81.676    23.131   .2960 Total Temperature:Heated #1 (oC)	TTHR1 	 -89.413    25.436  -.0021

Total Temperature:Heated #2 (C)      TTHR2   -91.553    25.993  -.0168

Vertical Press. Diff. Radome(mb)	ADIFR	   0.    0.9978  .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.204     0.9979  .0

Vertical Press. Diff. Gustpod(mb)	ADIFR_GP  0.0      -1.0     .0

Horizontal Press. Diff. Radome(mb) 	BDIFR_GP  0.0      -1.0     .0

Aircraft Pitch Angle (deg)	PITCH+	   0.35      1.0     .0

Static Pressure-Fuselage (hPa)	PSF	  -0.445     1.0006  .0

Dew Point Sensor Left Serial (C)	DPLS	  -0.046     1.029   .0

Dew Point Sensor Right Serial (C)	DPRS	  -0.046     1.029   .0

# +  Corrected for systematic offset

# Calibration Summary

page 2

Post Calibration – 9/24/09

Param

Variable                    Name      C1        C2       C3__

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.391	  17.232  -.0039

Dynamic Pressure-Radome (hPa)	QCR 	  -1.626	  17.304	-.0051

Total Temperature-Radome Left (oC)	TTFR	 -81.447    23.301   .2921 Total Temperature:Heated #1 (oC)	TTHR1 	 -89.121    25.616  -.0035

Total Temperature:Heated #2 (C)      TTHR2   -91.471    26.202  -.0242

Vertical Press. Diff. Radome(mb)	ADIFR	   0.234     0.9978  .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.204     0.9979  .0

Vertical Press. Diff. Gustpod(mb)	ADIFR_GP  0.0      -1.0     .0

Horizontal Press. Diff. Radome(mb) 	BDIFR_GP  0.0      -1.0     .0

Aircraft Pitch Angle (deg)	PITCH+	   0.43      1.0     .0

Static Pressure-Fuselage (hPa)	PSF	  -0.342     1.0007  .0

Dew Point Sensor Left Serial (C)	DPLS	   0.0       1.0     .0

Dew Point Sensor Right Serial (C)	DPRS	   0.0       1.0     .0

# Post cal used in final processing.

Calibration Maneuvers –

RAF uses sets of in-flight maneuvers specifically designed to help evaluate the performance of the 3-D wind measurement system on the aircraft.  Sets of maneuvers were flown on one test flight and one research flight: TF01, TF03 and FF02 respectively. In each case the variations observed fell within acceptable limits.

PMS - 1-D Probe Bin Sizes

The CDP probe was used on the project.  The minimum droplet size (microns) for each sampling bin are listed below:

CDP**

bin#1 		8.5 um

bin#2		 	9.52

bin#3			10.54

bin#4			11.56

bin#5			12.58

bin#6			13.6

bin#7			14.62

bin#8			15.64

bin#9			16.66

bin#10		17.68

bin#11		18.7

bin#12		19.75

bin#13		22.5

bin#14		25.84

bin#15		29.18

bin#16		32.52

bin#17		35.86

bin#18		39.2

bin#19		42.54

bin#20		45.88

bin#21		49.22

bin#22		52.56

bin#23		55.9

bin#24		59.24

bin#25		62.58

bin#26		65.92

bin#27		69.26

bin#28		72.6

bin#29		75.94

bin#30		79.28

bin#31		82.62