## Slide 1: AMS Performance at DC3-Test
Some preliminary results
Pedro Campuzano Jost, Douglas Day, Michael Lechner, Jose Luis Jimenez

## Slide 2: Overview
Instrument was in working order for all six flights. Data was acquired for >95% of flight time (the rest being filters and software mode switches).
Real time data was exported for all six flights to the G-V data system. After some minor software fixes, starting with RF03 the stream was fully functional
Topics
Inlet characterization: Still loose ends, but basic performance seems acceptable
UHSAS comparison
Stratospheric/UT aerosol measurements in clean air and after deep convection: AMS detection limit
Sampling of polluted air thru clouds: AMS time resolution
Size and chemical characterization of organics: brief case study

## Slide 3: Inlet
A new secondary diffuser was build for improved sampling of larger particles. At the same time, small modifications were performed on the pressure controlled inlet (PCI) system to make it work all the way to 50.000 ft
The transmission of the PCI showed little changes in transmission over a wide range of inlet flows, being close to 100% for submicron aerosol
To characterize the transmission of the inlet system up to the AMS, the following was investigated:
  Flow into the system under “free flow” conditions. Is the diffuser working? Is the pressure drop sufficient despite the filters to run the inlet?
  Comparison of concentrations and size distributions for free flow and “choked” flow thru rapid switches
  Comparison with calculated inlet transmissions
  Comparison with the UHSAS (RF02 and 03)

## Slide 4: Aircraft Speed vs Inlet Velocity, “free flow” conditions
RF01
RF02
RF03
RF06
RF05
RF04
Red: Scaled Airspeed
Black: Cabin flow speed
Green: Corrected inlet speed, free flow
Blue: Scaled airspeed x 1.4

## Slide 5: Aircraft Speed/Inlet Velocity vs Height, “free flow” conditions (actual ratios)
RF01
RF02
RF03
RF06
RF05
RF04
Red Line: “Ideal ratio”. Bottom range: 5 -45 kft

## Slide 6: Varying Inlet speeds (free, /2, /4) during RF05: Nitrate

## Slide 7: Varying Inlet speeds (free, /2, /4) during RF05: Sulfate
Only for low to midheights! Testing during RF06 for >40 kft showed the same (s.b.)

## Slide 8: Chemically Resolved Size Distributions during the switching

## Slide 9: Size Distributions during the switch (averages)

## Slide 10: Calculated trasmission(line only)
“Free flow”
“Free Flow”/2

## Slide 11: UHSAS comparison, RF02
Filter

## Slide 12: UHSAS Comparison, RF03

## Slide 13: Altitude Profiles (RF01, RF03)

## Slide 14: Altitude Profiles (RF05, RF06)

## Slide 15: Stratosphere/UT aerosol loadings
This is a switched dataset!

## Slide 16: Typical AMS Detection Limits

## Slide 17: AMS Response to fast changing conditions Missed Approaches I: COS

## Slide 18: Missed Approaches I: COS (detail)

## Slide 19: Speeding thru clouds on the way to COS (10 kft)…

## Slide 20: Missed Approaches II: BJC
Cloud base higher than COS, hence no sharp transitions as seen there.
The “sulfate mystery spike” still shows up here, though

## Slide 21: Speciated Mass Size Distributions

## Slide 22: Bloomfield metro vs RM Nat Park

## Slide 23: Things to do for DC3-test…
CFD modeling of the inlet in collaboration with RAF (Dave Rogers) would be very helpful to understand the exact behavior at high altitude and under high pitch and roll angles (would be extremely useful for all aerosol instruments in the HIAPER cabin).
Explore HR analysis and repeat UHSAS comparison according to Middlebrook et al, 2011
Take a closer look at the mass spectral signature of the organic aerosol measured in the stratosphere during RF06
Once available, explore correlations between AMS loadings with the other chemical data acquired during DC-3
Try to increase the data quality of the size distributions by developing an HR analysis procedure for them

## Slide 24: Things to do before the next campaign (RHINO 2014?)
Depending on funding integrate cryopump to reduce pre-flight pumping requirements and increase sensitivity.
Automate filter sampling.
Redesign inlet assembly for easier service
Add flow meter to PCI bypass flow line
Explore software options for remote control of the AMS
Fly with a heated HIMIL and temperature monitor next time.