APPENDIX B.7

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 1 of 4)

DO NOT WRITE IN THIS BLOCK

Project: 		___ICE-T_________		Installation Period:	_1 – 17 Jun 2011___

Aircraft: 		___C-130_________		Beginning Date:           __20 June 2011___

Instrument Number:	_________________		Ending Date: 		___3 August 2011__

1. 	Instrument: 		___Colorado State University CFDC-1H_____________________

2.	Function: 		___Ice Nuclei Detector___________________________________

3.	Principal Investigator:	___Paul J DeMott_______________________________________

Address: 		___Atmos. Sci. Dept, Colorado State University_______________

Telephone: 		___(970) -491 – 8257____________________________________

4. 	Instrument Operator(s): _Paul De Mott, Gavin McMeeking, Ryan Sullivan____________

5.	Is this instrument commercially produced? 	____No_____________________________

6.	If so, please list name and address of manufacturer:

________________________________________________________________________

7. 	Please list serial number of the instrument:

________________________________________________________________________

Please attach a copy of the manufacturer’s instruction manual for the device.  If this is not possible, attach a copy of those pages of the instruction manual which describe the principles of operation, hazard warnings, safety features, and safety rules.

8. 	If the instrument is not commercially produced, please provide information requested below:

Designed by: 	___Anthony Prenni______________________________________

Organization: 	___Colorado State University______________________________

Address: 		___ Atmos. Sci. Dept, Colorado State University ______________

Telephone: 		___(970) – 491 - 8414____________________________________

Built by: 	            ___Anthony Prenni, Paul DeMott___________________________

Organization:         ___CSU_______________________________________________

Address: 	            ___same______________________________________________

###### APPENDIX B.7

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 2 of 4)

9.	Describe principles of operation, hazard warnings, safety features:

_Water vapor mass transfer from thermally controlled ice walls sets aerosol processing__

_condition for ice nuclei activation in buoyancy modified vertical flow.  Pressurized____

_ refrigeration system used.  Uses N2 gas at times.  All motors covered.  No eye exposure

_to OPC laser.  Circuit breakers used for power surges____________________________

10. 	If the instrument is commercially produced, has it been modified? 	__N/A____________

11. 	If modified, describe the modification.

________________________________________________________________________

________________________________________________________________________

________________________________________________________________________

All investigators please answer the following:

12. 	Does the instrument contain, use, or produce:

Radioactive materials		__N___	Compressed gases		__Y___

Other ionizing radiation	__N___	Non-ionizing radiation	__N___

Flammable liquids		__N___	Laser				__Y___

Radar				__N___	Flammable gases		__N___

Explosive materials		__N___	Toxic materials		__N___

13. 	If any of the categories were checked, specify the material below (for example, amount, energy levels, physical form, etc.).

__150 L compressed gas cylinder of N2.  0.38 lbs R404A refrigerant in each of 2 ______

__compressors (120 – 200 Psig maximum)_____________________________________

14. 	Please list all other chemicals you will use on board this aircraft in your experiment.

___Silica gel, Molecular sieves, water_________________________________________

________________________________________________________________________

15.	If your experiment consumes or discharges materials, will you need to carry additional materials on board?

____________________No_________________________________________________

###### APPENDIX B.7

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 3 of 4)

16. 	 What and how much extra materials will you need to carry?

________________________________________________________________________________________________________________________________________________________________________________________________________________________

17. 	What kind of container will you need to carry these materials?

________________________________________________________________________

18. 	If the device utilizes a laser, please classify the laser according to ANSI Z 136.1-1973 (circle one).

Class:	I     II     III     IV

- If your laser will be operating at a wavelength that is not eye safe,  what procedures will

be established to minimize the danger to yourself and other project participants?

______laser is completely contained__________________________________________

20.	If you are using compressed gas cylinders, what is the maximum pressure expected for each cylinder type?

_____________2000 PSI of N2_____________________________________________

21.	Will you be re-filling any compressed gas cylinders yourself, either at JeffCO or during the field deployment?

_____________________No________________________________________________

22. 	Are there any other hazards associated with the instrument itself, the required ground support equipment or the experiment which have not so far been covered in this questionnaire?

_____________________No________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________

23. 	How would you describe the probability of an accident resulting from the presence and use of your instrument on board the NCAR aircraft?

__Extremely low.  Successfully operated for 17 years without a failure_______________

________________________________________________________________________

###### APPENDIX B.7

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 4 of 4)

24.	How would you describe the severity of such an accident?

__ Mild.  There are possible hazards from a complete loss of refrigerant from  a __ __compressor.  Refrigerant in this case occupies 40L and due to density (4X air)__

__would limit oxygen in this volume.  Chance for skin freezing if positioned at __

__point of a refrigerant leak.___________________________________________

25. 	What precautions will you take to decrease the probability and the severity of an accident?  If any documented safety procedures from your home facility or university are available, please attach a copy of said materials to this form.

_Follow RAF aircraft guidelines for integration of User equipment on the C-130.____

_Cover all motors to prevent fragmentation.  Stow refrigerant lines within confines __

_ of the rack.__________________________________________________________

_6/16/2011______		________________________________________________

Date				  Signature of principal investigator or operator

________________________________________________

Printed name of principal investigator or operator

________________________________________________

Reviewed by

________________________________________________

Date