APPENDIX B.10

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 1 of 4)

DO NOT WRITE IN THIS BLOCK

Project: 		___NOMADSS___		Installation Period:      8 Apr – 22 May 2013

Aircraft: 		____C-130_______		Beginning Date: 	__1 Jun 2013______

Instrument Number:	_________________		Ending Date: 		__15 Jul 2013_____

1. 	Instrument: 		_UW detector for oxidized Hg species (DOHGS)______________

2.	Function: 		_measures elemental and oxidized mercury (Hg)_______________

3.	Principal Investigator:	_Daniel Jaffe___________________________________________

Address: 	_University of Washington-Bothell, 18115 Campus Way NE, Bothell, WA 98011______________________________________

Telephone: 		_(425) 352-5357________________________________________

4. 	Instrument Operator(s): _Jesse L. Ambrose and Lynne Gratz_______________________

5.	Is this instrument commercially produced? 	_No________________________________

6.	If so, please list name and address of manufacturer: _N/A_________________________

7. 	Please list serial number of the instrument:

_N/A___________________________________________________________________

Please attach a copy of the manufacturer’s instruction manual for the device.  If this is not possible, attach a copy of those pages of the instruction manual which describe the principles of operation, hazard warnings, safety features, and safety rules.

8. 	If the instrument is not commercially produced, please provide information requested below:

Designed by: 	_Jesse L. Ambrose, Seth N. Lyman*________________________

Organization: 	_University of Washington-Bothell_________________________

Address: 		_18115 Campus Way NE, Bothell, WA 98011________________

Telephone: 		_(425) 352-3533________________________________________

Built by: 	            _Jesse L. Ambrose, Seth N. Lyman*________________________

Organization:         _Same as above_________________________________________

Address: 	            _Same as above_________________________________________

*Now at Energy Dynamics Laboratory, Eastern Utah Operations, Utah State University Research Foundation, Vernal Utah, 84078

###### APPENDIX B.10

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 2 of 4)

9.	Describe principles of operation, hazard warnings, safety features:

This instrument measures gas-phase elemental and total mercury, and it quantifies oxidized mercury as the difference between these.  The core of the instrument is two modified mercury vapor analyzers (Tekran, Inc., Model 2537B).  These analyzers collect mercury on gold bead traps, then heat the traps to desorb the collected mercury into a UV fluorescence analytical cell.  Sample air is preconditioned in two separate channels before reaching the analyzers.  The first channel passes air through a quartz glass tube filled with quartz wool and heated to 650 °C.  The second channel passes air through a quartz wool-filled PFA Teflon trap at room temperature.  The instrument also contains a calibration box that utilizes a mercury permeation tube.

Hazards: 1) The instrument uses compressed argon and nitrogen, which are asphyxiants.

2) The instrument contains three separate vials that each contains 0.25 mL of liquid elemental mercury.  These vials are used for calibration checks.  Mercury is toxic.

3) The instrument contains two ½ in. diameter by 6 in. length PFA tubes filled with soda lime, which is caustic.

10. 	If the instrument is commercially produced, has it been modified? 	_N/A_____________

11. 	If modified, describe the modification.  _The Tekran mercury analyzers have been modified to output communication signals to other components of the instrument_____________

All investigators please answer the following:

12. 	Does the instrument contain, use, or produce:

Radioactive materials		_No_____	Compressed gases		_Yes_____

Other ionizing radiation	_No_____	Non-ionizing radiation	_No_____

###### APPENDIX B.10

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 3 of 4)

Flammable liquids		_No_____	Laser				_No_____

Radar				_No_____	Flammable gases		_No_____

Explosive materials		_No_____	Toxic materials		_Yes_____

13. 	If any of the categories were checked, specify the material below (for example, amount, energy levels, physical form, etc.).

1) One argon (Ar) cylinder compressed at 2500 psi; one nitrogen (N2) cylinder at 2500 psi.

2) Three sealed PFA vials containing 0.25 mL liquid elemental mercury each

14. 	Please list all other chemicals you will use on board this aircraft in your experiment.

The instrument contains two ½ in. diameter by 6 in. length PFA tubes filled with soda lime, and two of these traps will be kept on board as spares.

15.	If your experiment consumes or discharges materials, will you need to carry additional materials on board?

N/A

16. 	 What and how much extra materials will you need to carry?

Two extra soda-lime traps

17. 	What kind of container will you need to carry these materials?

_Double-Zip-Lock bags in a polycarbonate box_________________________________

18. 	If the device utilizes a laser, please classify the laser according to ANSI Z 136.1-1973 (circle one). N/A

Class:	I     II     III     IV

- If your laser will be operating at a wavelength that is not eye safe,  what procedures will    be established to minimize the danger to yourself and other project participants?

_N/A___________________________________________________________________

20.	If you are using compressed gas cylinders, what is the maximum pressure expected for each cylinder type?

_2500 psi________________________________________________________________

21.	Will you be re-filling any compressed gas cylinders yourself, either at JeffCO or during the field deployment?

_No____________________________________________________________________

###### APPENDIX B.10

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 4 of 4)

22. 	Are there any other hazards associated with the instrument itself, the required ground support equipment or the experiment which have not so far been covered in this questionnaire?

No

23. 	How would you describe the probability of an accident resulting from the presence and use of your instrument on board the NCAR aircraft?

Very low

24.	How would you describe the severity of such an accident?

Likely not to be severe

25. 	What precautions will you take to decrease the probability and the severity of an accident?  If any documented safety procedures from your home facility or university are available, please attach a copy of said materials to this form.

A mercury spill kit will be available

_March 21, 2013____		________________________________________________

Date				  Signature of principal investigator or operator

________________________________________________

Printed name of principal investigator or operator

________________________________________________

Reviewed by

________________________________________________

Date