APPENDIX B.11

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 1 of 4)

DO NOT WRITE IN THIS BLOCK

Project: 		___NOMADSS___		Installation Period:      8 Apr – 22 May 2013

Aircraft: 		____C-130_______		Beginning Date: 	__1 Jun 2013______

Instrument Number:	_________________		Ending Date: 		__15 Jul 2013_____

1. 	Instrument: 		ACD PTRMS- Quadrupole

2.	Function: 		VOC flux Instrument

3.	Principal Investigator:	 Alex Guenther

Address: 		NCAR-ACD

Telephone: 		x1447

4. 	Instrument Operator(s):  Thomas Karl, Lisa Kasser, Stephen Shertz (NCAR ACD)

Martin Graus, Bin Yuan (NOAA)

5.	Is this instrument commercially produced? 	Yes, but repackages for aircraft use

6.	If so, please list name and address of manufacturer: Ionicon. Innsbruck, Austria

7. 	Please list serial number of the instrument: N/A

Please attach a copy of the manufacturer’s instruction manual for the device.  If this is not possible, attach a copy of those pages of the instruction manual which describe the principles of operation, hazard warnings, safety features, and safety rules.

8. 	If the instrument is not commercially produced, please provide information requested below:

Designed by:

Organization: 	Repackaged by NCAR-ACD

Address: 		_________________________________________________

Telephone: 		_________________________________________________

Built by: 	            Stephen Shertz- x 1042

Organization:         NCAR-ACD

Address: 	            _________________________________________________

###### APPENDIX B.11

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 2 of 4)

9.	Describe principles of operation, hazard warnings, safety features:

Principle of operation- Chemical ionization Mass Spectrometer. Cold cathode lamp ionizes water and ambient air sample… injected into mass spectrometer.

Hazard warning- high voltages accessible inside instrument by opening access panels. Panels will not be opened in flight.

Safety Feature –No radioactive source for ionization, enclosure around high voltage.

10. 	If the instrument is commercially produced, has it been modified?  Yes

11. 	If modified, describe the modification.-Repackaged, Rewired, custom inlet for aircraft operations.

All investigators please answer the following:

12. 	Does the instrument contain, use, or produce:

Radioactive materials		__N___	Compressed gases		___Y__

Other ionizing radiation	__N___	Non-ionizing radiation	___N__

Flammable liquids		__N___	Laser				___N__

Radar				__N___	Flammable gases		___N__

Explosive materials		__N___	Toxic materials		___N__

13. 	If any of the categories were checked, specify the material below (for example, amount, energy levels, physical form, etc.).

Compressed gas standard-Non toxic level of VOC in Nitrogen…size CL

14. 	Please list all other chemicals you will use on board this aircraft in your experiment.

Water, VOC standard mentioned above, Platinum Catalyst in catalytic converter

15.	If your experiment consumes or discharges materials, will you need to carry additional materials on board?

N/A

###### APPENDIX B.11

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 3 of 4)

16. 	 What and how much extra materials will you need to carry?  N/A

17. 	What kind of container will you need to carry these materials? N/A

18. 	If the device utilizes a laser, please classify the laser according to ANSI Z 136.1-1973 (circle one).

Class:	I     II     III     IV

- If your laser will be operating at a wavelength that is not eye safe,  what procedures will

be established to minimize the danger to yourself and other project participants?

______please attached a separate document covering this question__________________

20.	If you are using compressed gas cylinders, what is the maximum pressure expected for each cylinder type?

2216 psig in tank…3-5 psig in line

21.	Will you be re-filling any compressed gas cylinders yourself, either at JeffCO or during the field deployment?   No

22. 	Are there any other hazards associated with the instrument itself, the required ground support equipment or the experiment which have not so far been covered in this questionnaire?

Will have second VOC standard similar to above for ground use.

23. 	How would you describe the probability of an accident resulting from the presence and use of your instrument on board the NCAR aircraft?

Instrument has no inherent risk above and beyond typical COTS laboratory instrument.  There is no known way to put the instrument in an inherently unsafe condition. There are 	exposed high voltage surfaces inside of the instrument enclosures.  There are safety placards/ guards to prevent accidental contact.

###### APPENDIX B.11

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 4 of 4)

24.	How would you describe the severity of such an accident?

Contact with high voltage (3000 volt- few milliamps of current)- Mild shock.

25. 	What precautions will you take to decrease the probability and the severity of an accident?  If any documented safety procedures from your home facility or university are available, please attach a copy of said materials to this form.

Operators manual- Still to be completed.

___3/29/13_____		________________________________________________

Date				  Signature of principal investigator or operator

Stephen Shertz, Project Engineer/Operator

Printed name of principal investigator or operator

________________________________________________

Reviewed by

________________________________________________

Date