APPENDIX B.4

RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 1 of 4)

DO NOT WRITE IN THIS BLOCK

Project: 		___NOMADSS___		Installation Period:      8 Apr – 22 May 2013

Aircraft: 		____C-130_______		Beginning Date: 	__1 Jun 2013______

Instrument Number:	_________________		Ending Date: 		__15 Jul 2013_____

1. 	Instrument: 		______NO-NO2_______________________________

2.	Function: 		Measure in situ NO & NO2 mixing ratios__________

3.	Principal Investigator:	___A. J. Weinheimer__________________________

Address: 		___NCAR__________________________________

Telephone: 		___303-497-1444____________________________

4. 	Instrument Operator(s):  Weinheimer, Knapp, Montzka, Flocke, Campos, Rooney

5.	Is this instrument commercially produced? 	__No.__________________________

6.	If so, please list name and address of manufacturer:

7. 	Please list serial number of the instrument:

_________NA_____________________________________________________

Please attach a copy of the manufacturer’s instruction manual for the device.  If this is not possible, attach a copy of those pages of the instruction manual which describe the principles of operation, hazard warnings, safety features, and safety rules.

8. 	If the instrument is not commercially produced, please provide information requested below:

Designed by:

Ridley, Grahek, Weinheimer, Knapp, Montzka, Flocke, Hendershot

Organization: 	NCAR__________________________________________

Address: 		Boulder, CO____________________________________

Telephone: 		303-497-1444___________________________________

Built by: 	            Same as above.____________

Organization:         NCAR_____________________________________

Address: 	            Boulder, CO_________________________________

###### APPENDIX B.4

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 2 of 4)

9.	Describe principles of operation, hazard warnings, safety features:

Measures in situ NO and NO2.  Pumps ambient air through the instrument at a flow of 2 slpm.  Add reagent ozone that is produced in real time to generate a chemiluminescent reaction between O3 and ambient NO.  Count photons from this light to measure ambient NO.  A second flow channel uses 400 nm light to convert NO2 to NO, which is subsequently detected to give a measurement of NO2.  Instrument is autonomous and programmed to detect different failure modes, such as pump failure.  Under such circumstances valves are closed and instrument shut down is initiated.  All valves are normally closed, so in event of power outage, valves will close and halt gas flows.  Instrument uses about 10 lbs of dry ice, per flight, as a coolant.  This is added in pre-flight, with no need to add during flight.

10. 	If the instrument is commercially produced, has it been modified? 	_NA_______

11. 	If modified, describe the modification.

All investigators please answer the following:

12. 	Does the instrument contain, use, or produce:

Radioactive materials		__No__	Compressed gases		_Yes___

Other ionizing radiation	__No__	Non-ionizing radiation	400 nm light_

Flammable liquids		__No__	Laser				_No___

Radar				__No__	Flammable gases		_No___

Explosive materials		__No__	Toxic materials		_No___

13. 	If any of the categories were checked, specify the material below (for example, amount, energy levels, physical form, etc.).

Compressed gases:

- zero air:  1 size 150 cylinder at 2000 psi

###### APPENDIX B.4

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 3 of 4)

- oxygen:  1 size 150 cylinder at 2000 psi

- NO-in-N2 at 3 ppmv:  1 size 30 cylinder at 2000 psi

Non-ionizing radiation:

Near-uv light at 400 nm to photolyze NO2 to NO.  Is contained under a shroud.

14. 	Please list all other chemicals you will use on board this aircraft in your experiment.

Dry ice is used to cool the photomultiplier tubes.

15.	If your experiment consumes or discharges materials, will you need to carry additional materials on board?

Materials are consumed, but none need to be added during flight.  So, no need to carry additional.

16. 	 What and how much extra materials will you need to carry?

17. 	What kind of container will you need to carry these materials?

________________________________________________________________________

18. 	If the device utilizes a laser, please classify the laser according to ANSI Z 136.1-1973 (circle one).

Class:	I     II     III     IV

- If your laser will be operating at a wavelength that is not eye safe,  what procedures will

be established to minimize the danger to yourself and other project participants?

______please attached a separate document covering this question__________________

20.	If you are using compressed gas cylinders, what is the maximum pressure expected for each cylinder type?

________2000 psi for each (zero air, oxygen, NO-in-N2 at 3 ppm____________

21.	Will you be re-filling any compressed gas cylinders yourself, either at JeffCO or during the field deployment?

_____________________No.________________________________________________

22. 	Are there any other hazards associated with the instrument itself, the required ground support equipment or the experiment which have not so far been covered in this questionnaire?

No

###### APPENDIX B.4

# RAF PROJECT SAFETY COMMITTEE

HAZARDOUS MATERIALS AND DEVICES

(Page 4 of 4)

23. 	How would you describe the probability of an accident resulting from the presence and use of your instrument on board the NCAR aircraft?

24.	How would you describe the severity of such an accident?

25. 	What precautions will you take to decrease the probability and the severity of an accident?  If any documented safety procedures from your home facility or university are available, please attach a copy of said materials to this form.

Autonomous code is programmed to close valves and shut down the instrument in the event of a pump failure.

On power outage, all solenoid valves will close, stopping gas flows.

____3 / 7 / 13____		___________________________________________

Date				  Signature of principal investigator or operator

___Andrew J. Weinheimer___________________

Printed name of principal investigator or operator

________________________________________________

Reviewed by

________________________________________________

Date