Summary of Calibration Factors

GV - Basic PREDICT Configuration

Montgomery et al

Pre-Calibration – 7/30/10

Param

Variable                    Name      C1        C2       C3

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.126	  0.9966   .0

Dynamic Pressure-Radome (hPa)	QCR 	   0.16	  1.0025	 .0

Total Temp: Rosemont Heated #1 (C)	TTFH1	 -96.230    28.509  -.3163

Total Temp: Rosemont Heated #2 (C)	TTFH2   -95.780    28.020  -.2937

Total Temp: HARCO Heated #1 (C)	TTHR1 	 -89.200    25.923  -.0794

Total Temp: HARCO Heated #2 (C)      TTHR2   -89.310    26.128  -.0902

Vertical Press. Diff. Radome(mb)	ADIFR	  -0.128     1.002   .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR     0.060     1.0014  .0

Vertical Press. Diff. Gustpod(mb)	ADIFR_GP  0.0      -1.0     .0

Horizontal Press. Diff. Radome(mb) 	BDIFR_GP  0.0       1.0     .0

Aircraft Pitch Angle (deg)	PITCH+	   0.30      1.0     .0

Static Pressure-Fuselage (hPa)	PSF	  -0.445     1.0006  .0

Dew Point Sensor Left Serial (C)	DPLS	   0.0       1.0     .0

Dew Point Sensor Right Serial (C)	DPRS	   0.0       1.0     .0

Cabin Pressure (hPa)	PCAB     -4.350   108.196  -.0150

# +  Corrected for systematic offset

# Calibration Summary

page 2

Post Calibration – 10/06/10

Param

Variable                    Name      C1        C2       C3__

Dynamic Pressure-Fuselage L. (hPa)	QCF  	   0.184	   1.001   .0

Dynamic Pressure-Radome (hPa)	QCR 	  -0.073	   1.001	 .0

Total Temp: Rosemont Heated #1 (oC)	TTFH1	 -81.447    23.301   .2921

Total Temp: Rosemont Heated #2 (C)	TTFH2

Total Temp: HARCO Heated #1 (oC)	TTHR1 	 -89.235    25.925  -.0889

Total Temp: HARCO Heated #2 (C)      TTHR2   -91.471    26.202  -.0242

Vertical Press. Diff. Radome(mb)	ADIFR	   0.009     1.034  .0

Horizontal Press. Diff. Gustpod(mb) 	BDIFR    -0.007     0.9991  .0

Vertical Press. Diff. Gustpod(mb)	ADIFR_GP  0.0      -1.0     .0

Horizontal Press. Diff. Radome(mb) 	BDIFR_GP  0.0       1.0     .0

Aircraft Pitch Angle (deg)	PITCH	   0.30      1.0     .0

Static Pressure-Fuselage (hPa)	PSF	  -0.282     1.0014  .0

Dew Point Sensor Left Serial (C)	DPLS	     not functional

Dew Point Sensor Right Serial (C)	DPRS	  -0.507     0.986   .0002 Cabin Pressure (hPa)	PCAB     -3.304   109.52   -.0201

CVI Sample Pressure (hPa)	CVPCN    -5.48    342.47    .0

# Post cal used in final processing.

Calibration Maneuvers –

RAF uses sets of in-flight maneuvers specifically designed to help evaluate the performance of the 3-D wind measurement system on the aircraft.  Sets of maneuvers were flown on one test flight and one research flight: TF01, TF03 and FF02 respectively. In each case the variations observed fell within acceptable limits.

PMS - 1-D Probe Bin Sizes

CDP probe 016 was used on the project.  The minimum droplet size (microns) for each sampling bin are listed below:

CDP

bin#1 		 2.0 um

bin#2		 	 3.0

bin#3			 4.0

bin#4			 5.0

bin#5			 6.0

bin#6			 7.0

bin#7			 8.0

bin#8			 9.0

bin#9			10.0

bin#10		11.0

bin#11		12.0

bin#12		13.0

bin#13		14.0

bin#14		16.0

bin#15		18.0

bin#16		20.0

bin#17		22.0

bin#18		24.0

bin#19		26.0

bin#20		28.0

bin#21		30.0

bin#22		32.0

bin#23		34.0

bin#24		36.0

bin#25		38.0

bin#26		40.0

bin#27		42.0

bin#28		44.0

bin#29		46.0

bin#30		48.0

bin#31		50.0