RF02

TECH: Cowan

Date 1/09/09

Billings, Mt to Anchorage, Ak

Pre-flight:

Fixed main entry door.  On initial power up noticed a smell in the forward cabin by racks L-1 and R-1, then noticed smoke in the cabin.  Secured mission power.  Found that the UCATS inlet heater had gotten to hot and discolored the protective coating.  If the heater is left on when power is first applied it will over heat (but not burn out) and cause this problem. If it is smelling in that area let the operator know so that they can turn it off.  Late turning on the Dew Point sensors.

It is very easy to get side tracked with a single tech and you have to watch the check list to make sure that things get done.

2016 Block out

2023 Take off

Noticed two PLWC gaps of about 1 min each, one from 2035 - 2037 and the other from 2041 - 2043

Power readings:

Left Side A Bus

400 Hz Phase A 14 A (1A Jump) 115 V

400 Hz Phase B 12.5 (1A Jump) 116 V

400 Hz Phase C 14.5 (1A Jump) 115 V

400 Hz AI Phase A 0A

400 Hz AI Phase B 0A

400 Hz AI Phase C 0A

60 Hz 116V 20A

Right Side B Bus

400 Hz Phase A 8.5 (1A Jump) 115V

400 Hz Phase B 10.8 (1A Jump) 113V

400 Hz Phase C 8.5 (1A Jump) 116V

400 Hz AI Phase A 0A

400 Hz AI Phase B 4.0A 113V

400 Hz AI Phase C 0A

60 Hz 115V 12A

28V AI 29V 15A

Taken at about 1 hour into flight at 40K feet alt.

2125 PLWC went to max then flattened out then back to 12.  It took about 5 seconds to do that.  Seems to still be working.

2142 Dew points finally match. Waiting to see if they continue to function or if one needs to be reset.

2143 DPR started to react and now it is tracking with DPL. It still does not look like it is seeing really good structure.

2144 Large VCSEL spike at this time. Might be an equipment level shift.

On the first missed approach both Dew Points react there are some spikes.

2220 Looks like VCSEL spikes are becoming regular at about every 6 minutes.

2223 Looking back at the last 10 min of PLWC data seeing some large spikes up to 24 W. It always seems to come back and appears to be working.

2239 – 2251 lots of VCSEL spikes including a 3 minute place where it was flatlined  high.  Seems to be just a VCSEL issue.

2350 Looking at the Novatel cable closest to the aisle to make sure that it is plugged in solid.  It seems to be.

0018 to 0030 seeing PLWC spike in the 1 to 2 W range.  Might be real data but not seeing it in the other probes.  PLWC seems to be spiky this  flight.

0139 in cloud and have been for a while, CDP not seeing any data.  Might be because it is more snow.

0124 – 0136 lost all Garmin data for this time.

0151 Picked up some ice on the last approach.

0228 RW DSM stopped attempting reboot.

0251 cycle power on RW DSM trying to get it to come up again.

0307 Reboot did not help.  DSM will not start.

0404 Landing