RF04

TECH: Cowan

04/28/08

Pre-flight no NO-NOy displays this flight.

1510 Block out

1519 Takeoff

1530 PLWC stuck at 50W cycling AI power to see if it clears the problem.

1531 PLWC Power off but still seeing 50W on display.

1532 When power re-applied to instrument the condition cleared.

1535 Had another PLWC hang up , returned to 2W instead of 0w.

1547 High levels on PLWC again, cycling power to see if the above is repeatable.

This time the PLWC dropped down to a hard 0.

1549 Instrument back on at 2.2W unsure of what is going on and what it causing the spikes and level shifts.

1615 Looks like LWO Pod is reporting NOPPS.

Also noticed little spikes in data (less than 1 W).  Did not notice them before because of the big spikes. Suspect that the analog card is the issue.

1643 Seeing spikes in PLWC, Cycling power on DSM to correct.

1645 LWO DSM back online.

1707 Lost chat in a steep turn.

1711 Good 2DC data. Should be good data on all probes.

1856 Power Readings:

Left A Bus

400 A Phase 20.9 - 14.5 A

400 B Phase 12.6  A

400 C Phase 16.5 A

400  Phase 115 A Volts

400  Phase 117 B Volts

400  Phase 114 C Volts

400 AI

All 0 with 115V

60Hz  116 V  25 A

28Volts AI  29.1 V 9 A

Right Side B Bus

400 A Phase 8 A

400 B Phase 13 A

400 C Phase 7.2 A

400  Phase A 115 Volts

400  Phase B 114 Volts

400  Phase C 115 Volts

400 AI

A 0

B 3.8A 113V

C 0

60Hz 115 V 13 A

1840-1910 Dew Points looked weird. Unknown what they did. Both Dew Points reacted together.

1922 LWO DSM Messages:

Bad PPS – SYNC, YEAR, MAJTM, NOPPS (red), Code sync Status=0x4, IRIG-UNIX=0.05Sec

Good PPS - SYNC, YEAR, MAJTM, PPS, Code sync Status=0x0, IRIG-UNIX=0.05Sec

2025 Data on both PLWC and RICE.

2136 Both CDP and 2DC data also some RICE and maybe a little PLWC.

2216 Landing

On shutdown, data system did not halt correctly. Message was:

E100:eth3 e1000_watchdog_task:NIC Link is up 10 MBPS ½ duplex

ADDRCONF (netdev_change):eth3:link becomes ready

ADDRCONF(netdev_Up):eth3link is not ready

The message continues to repeat.

System was powered off in this state.