RF05

TECH:  G.B.

10/24/08

Takeoff:  06:32

06:35 – Power on the plug of Right Front 115V #2 does not seem to work.  Move plug to another outlet and checked on maintenance day.  Checked out ok and moved it back.

06:46 – CNCON is negative, but the CN Counter has counts.  Found that the 28V had been accidentally turned off.  Turned it back on and it works fine.

10:00 – 11:00 - See gap of nimbus to dsm server time stretch from 2 sec to 1:40.  Turned off the track plot on rear left rack display computer.  This solved the problem and nimbus corrected itself in about 10:00

12:24 – DPB completely fails.  Reads -102 on display with no service mirror or rebalance lights.  Try to all positions on selector knob with no result.  Find later that Q14 chip is blown on board of control unit.

13:00 – 13:30 – XUVI is somewhat flaky.

*Have seen a few spikes in OAT

GPS drops – 08:05, 10:33, 10:36, ?

Landing:  15:24