RF06

TECH: Cowan

05/01/08

On pre-flight 1404 DSM 312 (TDL) found to be stopped.  Restarted the DSM with the switch at the ADS station.

1948 Block out

1950 Take off

1955 VCSEL temperature and dew point are very weird numbers, and do not compare with our data at all.

1958 Saw a small amount of liquid water on climb out but not much.

2032 Bose headset have ringing. Both the R-3 and L-3 seats have reported this issue in the last two flights.  Originally thought that the issue was due to the window heat in those positions. Found out later that it was not the window heat and that the interference could be heard in any location.  It appears to be some kind of grounding issue since you can ground yourself to the rack and the interference will stop. Headsets need to be sent out for repair.

2103 Seeing smaller PLWC spikes at the .5 to 1W range.  I have not seen the larger spikes that we saw on previous flights.

2129 Large 50W peak in PLWC.

2135 PLWC and rice comparison at this time.

2140 PLWC 0W .

2144 Cycling power on PLWC instrument to correct 0W issue. It came back after this to 2W. This appears to be an instrument issue not an A/D issue.

2154 Still small spikes in PLWC.

2157 PLWC spiking DOES NOT correspond to aircraft radio usage, 2 spikes no radio comms.

2218 Cycle power on all LWO instruments to see if PLWC spiking will stop.

2224 Another set of small PLWC spikes. Resetting the DSM did not cure the problem.

2228 Other than the above spikes it seems to have stopped spiking do much.

2238 Numerous spikes in PLWC again.

2254 Power Readings:

Left A Bus

400 A Phase 13.5 – 18.8 A

400 B Phase  12 A

400 C Phase 15 A

400  Phase 115 A Volts

400  Phase 116 B Volts

400  Phase 114 C Volts

400 AI

All 0 with 115V

60Hz  116 V  20 A

28Volts AI  29.1 V 10 A

Right Side B Bus

400 A Phase 8.5 A

400 B Phase 13.8 A

400 C Phase 10.5  A

400  Phase A 115 Volts

400  Phase B 113 Volts

400  Phase C 115 Volts

400 AI

A 0

B 3.9A 113V

C 0

60Hz 115 V 12 A

2346 Seeing data on both the PLWC and the RICE.

0010 Landing