RF13

TECH: Cowan

06/16/08

RR out of Jeffco via Gulf of Mexico

1628 Block out

1706 Large spike in PLWC up to 40W. Also numerous smaller spikes also.

1806 Left DSM (DSM312) stopped, rebooted. Stopped at 1803 initially.

1820 Lost all chat ground connections. Waiting to see if it comes back.

1826 Chat reconnected.

1859 Went through some clouds. Saw data on both RICE and PLWC.

1908 After Low approach, both Dew points have SERVICE MIRROR lights blinking, REBALANCE is on solid.  Also Mouse and Keyboard locked up. Seeing -32767 at an ascii feed.  Cycled power on the Dew Pointers.

1910 Dew Pointer back on line.  Lost data from 1904 to 1910.

1930 to 1934 saw some spiking on PLWC but it seems to have quit.

1958 Dew Point again when entering the Marine Boundary Layer had SERVICE MIRROR and REBALANCE lights lit. Cycled Power.

2024 Right Dew Point, still has REBALANCE light on and seem to be lagging badly.

2035 Power reading:

Left A Bus

400 A Phase  13.5 -19.1A

400 B Phase  11.7 – 12.5 A

400 C Phase 14 A

400  Phase 115 A Volts

400  Phase 117 B Volts

400  Phase 115 C Volts

400 AI

All 0 with 115V A and C and B at 117V

60Hz  116 V  21 A

28Volts AI  29.1 V 10 A

Right Side B Bus

400 A Phase 8.3 A

400 B Phase 13.6 A

400 C Phase 8.8  A

400  Phase A 115 Volts

400  Phase B 114 Volts

400  Phase C 115 Volts

400 AI

A 0A 115V

B 3.9A 113V

C 0A 115V

Right Dew point REBALANCE Lit but so is DEW POINT light, so I am letting it stay that way at this point.

- 2156 Right DEWPOINT agrees with the VCSEL but the left is way off.  Will reboot the Left after the missed approach.

2159 Left Side Dew point, now with SERVICE MIRROR.  Leaving it to see if it will clear up once we start to climb. Display head is reading correct dewpoint but the acsii shows it as disconnected. Waiting to see if it will reconnect.

2202 Right Dew Point reading 35 to 40C and that seems high, waiting to see if it will recover when we climb.

2204 Looks like the Right Dew point has recovered abut the left is still bad.  Waiting to see if it will recover on it’s own.

2206 Reset Left Dew point since it had not recovered and we are above the altitude where it failed.

2211 Left Dew point working well and tracking with the VCSEL.  Right is not tracking, rebooting the RIGHT Dew point.

2214 Right Dew point seeing more structure but higher than Left.

2230 Right Dew point not depressing as far as the left at altitude.

0005 Landing

0006 Block in.